## Final
|  | COND | CONDF | RPD | Comments | Action | Result | pH | pHF | Difference | Comments | Action | Result
Station | Date | µmho/cm | µmho/cm | %
A30 | 2010-02-04 11:00 | 275 | 250 | 9.5238 |  |  |  | 6.9 | 6.2 | 0.7
A30 | 2010-02-24 11:00 | 224 | 188 | 17.4757 |  |  |  | 7 | 6.6 | 0.4
A30 | 2010-03-02 13:20 | 213 | 209 | 1.8957 |  |  |  | 7.1 | 6.82 | 0.28
A30 | 2010-03-28 12:30 | 201 | 160 | 22.7147 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.3 | 6.87 | 0.43
A30 | 2010-04-10 16:30 | 202 | 160 | 23.2044 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on April 11, 2010 suggests calibration on field meter may have been off. | 7.3 | 6.76 | 0.54
A30 | 2010-04-19 11:55 | 187 | 150 | 21.9585 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.3 | 6.98 | 0.32
A30 | 2010-05-01 16:00 | 1670 | 1610 | 3.6585 |  |  |  | 3.6 | 3.77 | 0.17
A30 | 2010-05-18 14:35 | 752 | 132 | 140.2715 | Field and lab values correctly entered into emLine. | Request Retest | Retest not possible because was requested too late. Remainder of samples had been discarded. Discrepancy between values remains. | 4.9 | 5.22 | 0.32
A30 | 2010-06-02 12:07 | 484 | 549 | 12.5847 |  |  |  | 5 | 5.65 | 0.65
A30 | 2010-06-15 12:22 | 320 | 296 | 7.7922 |  |  |  | 6.7 | 6.16 | 0.54
A30 | 2010-07-07 14:35 | 753 | 810 | 7.2937 |  |  |  | 5.82 | 6.09 | 0.27
A30 | 2010-07-22 13:30 | 410 | 371 | 9.9872 |  |  |  | 6.53 | 7 | 0.47
A30 | 2010-08-05 12:15 | 312 | 339 | 8.2949 |  |  |  | 6.75 | 6.85 | 0.1
A30 | 2010-08-18 09:30 | 238 | 260 | 8.8353 |  |  |  | 6.8 | 7.68 | 0.88
A30 | 2010-09-01 16:57 | 224 | 227 | 1.3304 |  |  |  | 6.99 | 6.69 | 0.3
A30 | 2010-10-20 11:20 | 223 | 208 | 6.9606 |  |  |  | 7.36 | 6.56 | 0.8
A30 | 2010-11-04 11:10 | 221 | 223 | 0.9009 |  |  |  | 7.4 | 7.33 | 0.07
A30 | 2010-12-01 12:05 | 197 | 80 | 84.4765 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. Multiple flagged conductivity values on December 1, 2010 suggests calibration on field meter may have been off. | 7.17 | 6.86 | 0.31
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
ETA Combined | 2010-09-02 08:34 | 7060 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 4.85 | 5.98 | 1.13 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
ETA Combined | 2010-10-20 10:40 | 7510 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 4.89 | 6.3 | 1.41 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
ETA Combined | 2010-11-08 10:55 | 7450 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.05 | 6.05 | 1 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
ETA Combined | 2010-12-01 09:25 | 7360 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.29 | 6.24 | 0.95
FAROCR | 2010-05-03 14:44 | 50 | 78 | 43.75 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.3 | 7.54 | 0.24
FAROCR | 2010-06-05 12:57 | 29 | 21 | 32 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.2 | 7.84 | 0.64
FAROCR | 2010-07-08 13:42 | 48 | 42 | 13.3333 |  |  |  | 7.5 | 8.16 | 0.66
FAROCR | 2010-08-03 12:30 | 81 | 79 | 2.5 |  |  |  | 7.79 | 7.33 | 0.46
FAROCR | 2010-09-01 14:40 | 85 | 80 | 6.0606 |  |  |  | 7.69 | 7.75 | 0.06
FAROCR | 2010-10-19 13:27 | 77 | 244 | 104.0498 | Field and lab values correctly entered into emLine. | Request Retest | Retest performed, see results below. | 7.67 | 7.26 | 0.41
FAROCR Retest | 2010-10-19 13:27 | 74 | 244 | 106.9182 | Retest for October 19 FAROCR EC. Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.67 | 7.26 | 0.41
FCO | 2010-06-16 15:52 | 118 | 122 | 3.3333 |  |  |  | 7.1 | 7.85 | 0.75
FCO | 2010-10-07 16:29 | 123 | 70 | 54.9223 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. | 7.2 | 7.52 | 0.32
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
FCS-4 | 2010-09-02 08:51 | 6080 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.28 | 6.34 | 1.06 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
FCS-4 | 2010-10-20 09:27 | 7180 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 4.61 | 6.5 | 1.89 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains.
FCS-4 | 2010-11-10 10:34 | 7160 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.06 | 6.63 | 1.57 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past sample hold time. Discrepancy between values remains.
FCS-4 | 2010-12-02 14:35 | 6890 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 4.73 | 6.67 | 1.94 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past sample hold time. Discrepancy between values remains.
GDHSECK | 2010-03-04 12:06 | 1230 | 1110 | 10.2564 |  |  |  | 8.1 | 7.56 | 0.54
GDHSECK | 2010-06-16 15:30 | 1230 | 1238 | 0.6483 |  |  |  | 8.1 | 8.19 | 0.09
GDHSECK | 2010-10-06 13:54 | 1110 | 1234 | 10.5802 |  |  |  | 8.07 | 7.05 | 1.02 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
NE1 | 2010-06-01 14:45 | 480 | 627 | 26.5583 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.8 | 7.46 | 0.34
NF1 | 2010-02-23 16:07 | 297 | 240 | 21.2291 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on February 23, 2010 suggests calibration on field meter may have been off. | 7.9 | 7.77 | 0.13
NF1 | 2010-03-10 13:21 | 298 | 269 | 10.2293 |  |  |  | 8 | 7.84 | 0.16
NF1 | 2010-04-11 15:30 | 325 | 260 | 22.2222 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on April 11, 2010 suggests calibration on field meter may have been off. | 7.9 | 7.37 | 0.53
NF1 | 2010-05-03 13:56 | 103 | 90 | 13.4715 |  |  |  | 7.6 | 7.43 | 0.17
NF1 | 2010-06-05 08:36 | 97 | 99 | 2.0408 |  |  |  | 7.6 | 7.84 | 0.24
NF1 | 2010-07-08 14:37 | 134 | 129 | 3.8023 |  |  |  | 7.91 | 7.71 | 0.2
NF1 | 2010-08-03 13:55 | 355 | 421 | 17.0103 |  |  |  | 7.26 | 6.68 | 0.58
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
NF2 | 2010-01-12 11:00 | 286 | 266 | 7.2464 |  |  |  | 7.7 | 6.86 | 0.84
NF2 | 2010-02-23 15:10 | 295 | 230 | 24.7619 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on February 23, 2010 suggests calibration on field meter may have been off. | 8 | 7.71 | 0.29
NF2 | 2010-03-10 10:41 | 298 | 267 | 10.9735 |  |  |  | 8 | 7.62 | 0.38
NF2 | 2010-04-14 10:26 | 308 | 270 | 13.1488 |  |  |  | 8.2 | 7.48 | 0.72
NF2 | 2010-05-03 16:50 | 107 | 97 | 9.8039 |  |  |  | 7.6 | 7.29 | 0.31
NF2 | 2010-06-04 15:18 | 87 | 81 | 7.1429 |  |  |  | 7.6 | 7.86 | 0.26
NF2 | 2010-07-08 11:00 | 136 | 154 | 12.4138 |  |  |  | 7.91 | 8.18 | 0.27
NF2 | 2010-08-03 15:30 | 196 | 193 | 1.5424 |  |  |  | 8.01 | 7.5 | 0.51
NF2 | 2010-09-01 09:56 | 203 | 222 | 8.9412 |  |  |  | 7.92 | 8.1 | 0.18
NF2 | 2010-10-18 14:03 | 215 | 200 | 7.2289 |  |  |  | 7.9 | 7.46 | 0.44
NF2 | 2010-11-09 13:33 | 258 | 252 | 2.3529 |  |  |  | 8.06 | 8.08 | 0.02
NF2 | 2010-12-01 17:15 | 262 | 120 | 74.3455 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. Multiple flagged conductivity values on December 1, 2010 suggests calibration on field meter may have been off. | 7.93 | 7.41 | 0.52
NFRC SC-1 | 2010-01-12 11:10 | 288 | 252 | 13.3333 |  |  |  | 7.6 | 6.84 | 0.76
NFRC SC-1 | 2010-02-23 14:46 | 298 | 250 | 17.5182 |  |  |  | 8 | 7.75 | 0.25
NFRC SC-1 | 2010-03-10 10:49 | 304 | 269 | 12.2164 |  |  |  | 8 | 7.83 | 0.17
NFRC SC-1 | 2010-04-14 11:00 | 304 | 260 | 15.6028 |  |  |  | 8.1 | 7.19 | 0.91
NFRC SC-1 | 2010-05-03 16:54 | 108 | 96 | 11.7647 |  |  |  | 7.6 | 7.57 | 0.03
NFRC SC-1 | 2010-06-04 15:40 | 88 | 84 | 4.6512 |  |  |  | 7.6 | 7.65 | 0.05
NFRC SC-1 | 2010-07-08 11:12 | 137 | 129 | 6.015 |  |  |  | 7.89 | 8.05 | 0.16
NFRC SC-1 | 2010-08-03 15:40 | 197 | 192 | 2.5707 |  |  |  | 8.04 | 7.64 | 0.4
NFRC SC-1 | 2010-10-18 14:10 | 217 | 197 | 9.6618 |  |  |  | 8.03 | 7.47 | 0.56
NFRC SC-2 | 2010-01-12 11:20 | 287 | 250 | 13.7803 |  |  |  | 7.6 | 6.95 | 0.65
NFRC SC-2 | 2010-02-23 14:15 | 300 | 240 | 22.2222 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.9 | 7.77 | 0.13
NFRC SC-2 | 2010-03-10 10:56 | 301 | 269 | 11.2281 |  |  |  | 8 | 7.61 | 0.39
NFRC SC-2 | 2010-04-14 12:05 | 309 | 270 | 13.4715 |  |  |  | 8.1 | 7.24 | 0.86
NFRC SC-2 | 2010-05-03 16:46 | 109 | 99 | 9.6154 |  |  |  | 7.6 | 7.72 | 0.12
NFRC SC-2 | 2010-06-04 16:04 | 89 | 86 | 3.4286 |  |  |  | 7.6 | 7.68 | 0.08
NFRC SC-2 | 2010-07-08 11:26 | 135 | 129 | 4.5455 |  |  |  | 7.82 | 7.97 | 0.15
NFRC SC-2 | 2010-08-03 15:55 | 196 | 186 | 5.2356 |  |  |  | 8.04 | 7.54 | 0.5
NFRC SC-2 | 2010-10-18 15:30 | 219 | 198 | 10.0719 |  |  |  | 8.03 | 7.47 | 0.56
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
NFRC SC-3 | 2010-01-12 10:50 | 288 | 257 | 11.3761 |  |  |  | 7.8 | 7 | 0.8
NFRC SC-3 | 2010-02-23 14:20 | 297 | 260 | 13.2855 |  |  |  | 8 | 7.66 | 0.34
NFRC SC-3 | 2010-03-10 11:14 | 304 | 270 | 11.8467 |  |  |  | 8 | 7.68 | 0.32
NFRC SC-3 | 2010-04-14 14:01 | 308 | 270 | 13.1488 |  |  |  | 8.1 | 7.31 | 0.79
NFRC SC-3 | 2010-05-03 16:58 | 109 | 96 | 12.6829 |  |  |  | 7.6 | 7.46 | 0.14
NFRC SC-3 | 2010-06-04 16:20 | 89 | 84 | 5.7803 |  |  |  | 7.6 | 7.69 | 0.09
NFRC SC-3 | 2010-07-08 11:43 | 136 | 130 | 4.5113 |  |  |  | 7.9 | 7.97 | 0.07
NFRC SC-3 | 2010-08-03 16:13 | 195 | 191 | 2.0725 |  |  |  | 8.03 | 7.69 | 0.34
NFRC SC-3 | 2010-10-18 15:10 | 218 | 197 | 10.1205 |  |  |  | 8.04 | 7.56 | 0.48
NFRC SC-4 | 2010-01-11 14:30 | 294 | 255 | 14.2077 |  |  |  | 7.8 | 6.95 | 0.85
NFRC SC-4 | 2010-02-23 13:16 | 303 | 260 | 15.2753 |  |  |  | 7.9 | 7.49 | 0.41
NFRC SC-4 | 2010-03-10 11:24 | 305 | 271 | 11.8056 |  |  |  | 8 | 7.82 | 0.18
NFRC SC-4 | 2010-04-14 14:38 | 316 | 270 | 15.6997 |  |  |  | 8.1 | 7.47 | 0.63
NFRC SC-4 | 2010-05-03 17:03 | 112 | 95 | 16.4251 |  |  |  | 7.7 | 7.54 | 0.16
NFRC SC-4 | 2010-06-04 16:30 | 89 | 86 | 3.4286 |  |  |  | 7.6 | 7.71 | 0.11
NFRC SC-4 | 2010-07-08 11:54 | 136 | 128 | 6.0606 |  |  |  | 7.83 | 7.87 | 0.04
NFRC SC-4 | 2010-08-03 16:20 | 196 | 203 | 3.5088 |  |  |  | 8.06 | 7.88 | 0.18
NFRC SC-4 | 2010-10-18 14:40 | 220 | 195 | 12.0482 |  |  |  | 8.13 | 7.58 | 0.55
NWID | 2010-02-22 13:33 | 399 | 340 | 15.9675 |  |  |  | 8.1 | 8.01 | 0.09
NWID | 2010-03-10 15:03 | 405 | 362 | 11.2125 |  |  |  | 8.1 | 7.87 | 0.23
NWID | 2010-04-15 10:22 | 411 | 420 | 2.1661 |  |  |  | 8.3 | 7.74 | 0.56
NWID | 2010-05-03 15:20 | 155 | 137 | 12.3288 |  |  |  | 7.7 | 7.77 | 0.07
NWID | 2010-06-04 18:00 | 202 | 208 | 2.9268 |  |  |  | 8 | 8.03 | 0.03
NWID | 2010-07-10 13:07 | 293 | 307 | 4.6667 |  |  |  | 8.05 | 8.15 | 0.1
NWID | 2010-08-03 12:33 | 336 | 340 | 1.1834 |  |  |  | 8.19 | 7.79 | 0.4
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
R10 | 2010-01-11 12:45 | 283 | 241 | 16.0305 |  |  |  | 7.7 | 7.06 | 0.64
R10 | 2010-02-23 10:42 | 292 | 235 | 21.6319 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on February 23, 2010 suggests calibration on field meter may have been off. | 8.1 | 7.87 | 0.23
R10 | 2010-03-10 13:35 | 292 | 259 | 11.9782 |  |  |  | 8 | 7.79 | 0.21
R10 | 2010-04-11 16:15 | 298 | 240 | 21.5613 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on April 11, 2010 suggests calibration on field meter may have been off. | 7.9 | 7.43 | 0.47
R10 | 2010-05-03 14:12 | 111 | 95 | 15.534 |  |  |  | 7.7 | 7.7 | 0
R10 | 2010-06-05 08:55 | 98 | 103 | 4.9751 |  |  |  | 7.7 | 7.93 | 0.23
R10 | 2010-07-08 14:25 | 132 | 125 | 5.4475 |  |  |  | 7.84 | 7.94 | 0.1
R10 | 2010-08-03 13:35 | 190 | 188 | 1.0582 |  |  |  | 8.11 | 7.46 | 0.65
R10 | 2010-09-01 15:35 | 193 | 187 | 3.1579 |  |  |  | 7.99 | 7.79 | 0.2
R10 | 2010-10-19 15:32 | 213 | 207 | 2.8571 |  |  |  | 7.95 | 7.65 | 0.3
R10 | 2010-11-10 16:56 | 256 | 234 | 8.9796 |  |  |  | 7.94 | 7.86 | 0.08
R10 | 2010-12-01 16:01 | 258 | 350 | 30.2632 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 8.03 | 7.58 | 0.45
R7 | 2010-01-11 11:45 | 254 | 235 | 7.771 |  |  |  | 7.9 | 7.21 | 0.69
R7 | 2010-02-23 12:08 | 266 | 260 | 2.2814 |  |  |  | 8 | 7.95 | 0.05
R7 | 2010-03-10 14:20 | 268 | 239 | 11.4398 |  |  |  | 8.1 | 7.78 | 0.32
R7 | 2010-04-11 13:30 | 275 | 240 | 13.5922 |  |  |  | 8.1 | 7.78 | 0.32
R7 | 2010-05-03 14:51 | 106 | 91 | 15.2284 |  |  |  | 7.8 | 7.66 | 0.14
R7 | 2010-06-05 13:20 | 101 | 97 | 4.0404 |  |  |  | 7.7 | 7.85 | 0.15
R7 | 2010-07-08 13:53 | 128 | 122 | 4.8 |  |  |  | 7.89 | 8.04 | 0.15
R7 | 2010-08-03 12:07 | 180 | 183 | 1.6529 |  |  |  | 8.13 | 7.37 | 0.76
R7 | 2010-09-01 14:28 | 180 | 186 | 3.2787 |  |  |  | 8.03 | 7.84 | 0.19
R7 | 2010-10-19 13:46 | 198 | 198 | 0 |  |  |  | 7.98 | 7.2 | 0.78
R7 | 2010-11-10 13:26 | 229 | 215 | 6.3063 |  |  |  | 8.02 | 7.9 | 0.12
R7 | 2010-12-01 14:05 | 239 | 100 | 82.0059 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. Multiple flagged conductivity values on December 1, 2010 suggests calibration on field meter may have been off. | 8.04 | 7.12 | 0.92
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
R8 | 2010-01-11 12:05 | 256 | 221 | 14.6751 |  |  |  | 7.9 | 6.91 | 0.99
R8 | 2010-02-23 11:54 | 252 | 220 | 13.5593 |  |  |  | 8.2 | 7.94 | 0.26
R8 | 2010-03-10 14:02 | 271 | 244 | 10.4854 |  |  |  | 8.1 | 7.84 | 0.26
R8 | 2010-04-11 14:00 | 274 | 230 | 17.4603 |  |  |  | 8.1 | 7.73 | 0.37
R8 | 2010-05-03 14:37 | 101 | 87 | 14.8936 |  |  |  | 7.7 | 7.51 | 0.19
R8 | 2010-06-05 12:35 | 89 | 89 | 0 |  |  |  | 7.7 | 7.92 | 0.22
R8 | 2010-07-08 13:29 | 123 | 121 | 1.6393 |  |  |  | 7.99 | 8.09 | 0.1
R8 | 2010-08-03 12:43 | 176 | 177 | 0.56657 |  |  |  | 8.13 | 7.26 | 0.87
R8 | 2010-09-01 14:53 | 172 | 172 | 0 |  |  |  | 8 | 7.66 | 0.34
R8 | 2010-10-19 14:24 | 191 | 193 | 1.0417 |  |  |  | 7.98 | 7.8 | 0.18
R8 | 2010-11-10 13:04 | 229 | 195 | 16.0377 |  |  |  | 8.03 | 7.76 | 0.27
R8 | 2010-12-01 14:55 | 233 | 100 | 79.8799 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. Multiple flagged conductivity values on December 1, 2010 suggests calibration on field meter may have been off. | 8.05 | 7.28 | 0.77
R9 | 2010-01-11 12:20 | 279 | 272 | 2.5408 |  |  |  | 7.8 | 7.11 | 0.69
R9 | 2010-02-23 11:30 | 289 | 250 | 14.4712 |  |  |  | 8.1 | 7.73 | 0.37
R9 | 2010-03-10 13:48 | 293 | 262 | 11.1712 |  |  |  | 8.1 | 7.88 | 0.22
R9 | 2010-04-11 17:00 | 292 | 230 | 23.7548 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. Multiple flagged conductivity values on April 11, 2010 suggests calibration on field meter may have been off. | 8.1 | 7.81 | 0.29
R9 | 2010-05-03 14:22 | 108 | 95 | 12.8079 |  |  |  | 7.7 | 7.78 | 0.08
R9 | 2010-06-05 12:00 | 97 | 101 | 4.0404 |  |  |  | 7.7 | 8.05 | 0.35
R9 | 2010-07-08 14:15 | 1 | 127 | 196.875 | Field and lab values correctly entered into emLine. | Request Retest | Retest not possible because was requested too late. Remainder of samples had been discarded. Discrepancy between values remains. | 7.91 | 8.01 | 0.1
R9 | 2010-08-03 13:12 | 190 | 194 | 2.0833 |  |  |  | 8.15 | 7.47 | 0.68
R9 | 2010-09-01 15:20 | 197 | 187 | 5.2083 |  |  |  | 7.88 | 7.81 | 0.07
R9 | 2010-10-19 14:57 | 213 | 211 | 0.9434 |  |  |  | 7.99 | 7.6 | 0.39
R9 | 2010-11-10 15:30 | 254 | 254 | 0 |  |  |  | 8.01 | 7.87 | 0.14
R9 | 2010-12-01 15:13 | 260 | 120 | 73.6842 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. Multiple flagged conductivity values on December 1, 2010 suggests calibration on field meter may have been off. | 7.86 | 7.46 | 0.4
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
RCSG#4 | 2010-01-12 11:40 | 313 | 310 | 0.96308 |  |  |  | 7.9 | 7.23 | 0.67
RCSG#4 | 2010-03-10 11:47 | 320 | 285 | 11.5702 |  |  |  | 8 | 7.75 | 0.25
RCSG#4 | 2010-04-13 15:12 | 331 | 300 | 9.8257 |  |  |  | 7.8 | 7.35 | 0.45
RCSG#4 | 2010-05-03 16:09 | 130 | 115 | 12.2449 |  |  |  | 7.8 | 7.64 | 0.16
RCSG#4 | 2010-06-04 11:15 | 92 | 92 | 0 |  |  |  | 7.6 | 7.84 | 0.24
RCSG#4 | 2010-07-08 15:47 | 126 | 126 | 0 |  |  |  | 7.78 | 7.94 | 0.16
RCSG#4 | 2010-08-03 15:20 | 200 | 207 | 3.4398 |  |  |  | 8 | 7.5 | 0.5
SP5-6 | 2010-06-01 15:36 | 1050 | 1172 | 10.9811 |  |  |  | 8 | 7.45 | 0.55
SP5-6 | 2010-07-10 13:43 | 1080 | 1049 | 2.9122 |  |  |  | 7.65 | 7.5 | 0.15
SP5-6 | 2010-08-05 10:40 | 1280 | 1216 | 5.1282 |  |  |  | 7.87 | 7.49 | 0.38
SP5-6 | 2010-09-01 11:57 | 1450 | 1439 | 0.76151 |  |  |  | 7.91 | 7.63 | 0.28
SP5-6 | 2010-10-06 15:32 | 1590 | 1786 | 11.6114 |  |  |  | 8.02 | 6.84 | 1.18 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
W10 | 2010-05-03 15:13 | 77 | 109 | 34.4086 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.7 | 7.71 | 0.01
W10 | 2010-06-05 15:23 | 75 | 68 | 9.7902 |  |  |  | 7.6 | 7.76 | 0.16
W10 | 2010-07-10 14:04 | 103 | 101 | 1.9608 |  |  |  | 7.67 | 8.08 | 0.41
W5 | 2010-06-01 15:55 | 752 | 1680 | 76.3158 | Field and lab values correctly entered into emLine. | Request Retest | Retest not possible because was requested too late. Remainder of samples had been discarded. Discrepancy between values remains. | 8.1 | 7.63 | 0.47
W8 | 2010-06-17 11:00 | 116 | 106 | 9.009 |  |  |  | 7.8 | 7.72 | 0.08
Weir 3 | 2010-03-04 10:47 | 1780 | 1730 | 2.849 |  |  |  | 7.8 | 6.97 | 0.83
Weir 3 | 2010-08-05 12:04 | 1840 | 1870 | 1.6173 |  |  |  | 7.61 | 6.96 | 0.65
X10 | 2010-01-11 13:05 | 326 | 286 | 13.0719 |  |  |  | 8 | 7.13 | 0.87
X10 | 2010-02-22 14:47 | 329 | 280 | 16.092 |  |  |  | 8 | 7.46 | 0.54
X10 | 2010-03-10 08:38 | 334 | 292 | 13.4185 |  |  |  | 8.1 | 7.51 | 0.59
X10 | 2010-04-13 16:08 | 340 | 310 | 9.2308 |  |  |  | 7.9 | 7.55 | 0.35
X10 | 2010-05-03 16:01 | 129 | 112 | 14.1079 |  |  |  | 7.8 | 7.79 | 0.01
X10 | 2010-06-04 09:43 | 99 | 111 | 11.4286 |  |  |  | 7.8 | 8.05 | 0.25
X10 | 2010-07-08 16:06 | 134 | 128 | 4.5802 |  |  |  | 7.95 | 8.09 | 0.14
X10 | 2010-08-03 13:30 | 205 | 205 | 0 |  |  |  | 8.18 | 8.01 | 0.17
X10 | 2010-09-02 11:34 | 196 | 194 | 1.0256 |  |  |  | 8.01 | 7.42 | 0.59
X10 | 2010-10-21 10:30 | 257 | 244 | 5.1896 |  |  |  | 8.16 | 7.93 | 0.23
X10 | 2010-11-09 16:13 | 274 | 268 | 2.214 |  |  |  | 8.18 | 7.52 | 0.66
X10 | 2010-12-01 17:40 | 275 | 240 | 13.5922 |  |  |  | 8.03 | 7.87 | 0.16
X11 | 2010-03-04 10:28 | 2650 | 2460 | 7.4364 |  |  |  | 7.6 | 7.05 | 0.55
X11 | 2010-08-05 11:46 | 2540 | 2289 | 10.3955 |  |  |  | 7.37 | 6.91 | 0.46
X12 | 2010-03-04 11:05 | 1560 | 1460 | 6.6225 |  |  |  | 7.8 | 7.27 | 0.53
X12 | 2010-08-05 12:05 | 1490 | 1243 | 18.0754 |  |  |  | 7.65 | 7.3 | 0.35
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X13 | 2010-01-11 12:53 | 1700 | 1430 | 17.2524 |  |  |  | 7.9 | 6.08 | 1.82 | Field and lab values correctly entered into emLine. | Request Retest | Retest not possible because was requested too late. Remainder of samples had been discarded. Discrepancy between values remains.
X13 | 2010-02-17 15:30 | 2220 | 2240 | 0.89686 |  |  |  | 7.8 | 7.17 | 0.63
X13 | 2010-03-02 11:50 | 2220 | 1780 | 22 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.7 | 7 | 0.7
X13 | 2010-04-15 11:06 | 2100 | 2220 | 5.5556 |  |  |  | 7.8 | 6.96 | 0.84
X13 | 2010-05-06 08:51 | 2240 | 2229 | 0.49228 |  |  |  | 7.6 | 7.13 | 0.47
X13 | 2010-06-10 10:58 | 2030 | 1859 | 8.794 |  |  |  | 8 | 7 | 1 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X13 | 2010-07-11 13:22 | 2160 | 1999 | 7.7422 |  |  |  | 7.49 | 7.38 | 0.11
X13 | 2010-08-05 11:31 | 2110 | 1798 | 15.9672 |  |  |  | 7.51 | 7.12 | 0.39
X13 | 2010-09-02 12:02 | 2050 | 2164 | 5.4105 |  |  |  | 7.6 | 7.2 | 0.4
X13 | 2010-10-07 14:35 | 2170 | 1270 | 52.3256 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. | 7.55 | 6.19 | 1.36 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X13 | 2010-11-04 12:11 | 2200 | 2165 | 1.6037 |  |  |  | 7.8 | 7.17 | 0.63
X13 | 2010-12-02 08:47 | 2310 | 2400 | 3.8217 |  |  |  | 7.57 | 7.07 | 0.5
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X14 | 2010-01-11 12:32 | 741 | 624 | 17.1429 |  |  |  | 8 | 6.76 | 1.24 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X14 | 2010-02-22 14:20 | 765 | 720 | 6.0606 |  |  |  | 8 | 6.91 | 1.09 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X14 | 2010-03-10 08:52 | 837 | 757 | 10.0376 |  |  |  | 7.9 | 7.71 | 0.19
X14 | 2010-03-22 17:46 | 894 | 749 | 17.6506 |  |  |  | 8.2 | 7.49 | 0.71
X14 | 2010-03-30 10:50 | 1130 | 1030 | 9.2593 |  |  |  | 8 | 7.37 | 0.63
X14 | 2010-04-06 08:24 | 1170 | 1100 | 6.1674 |  |  |  | 8 | 7.4 | 0.6
X14 | 2010-04-13 09:04 | 1160 | 1027 | 12.1628 |  |  |  | 8 | 7.85 | 0.15
X14 | 2010-04-20 14:43 | 873 | 1914 | 74.704 | Field and lab values correctly entered into emLine. | Request Retest | Retest not possible because was requested too late. Remainder of samples had been discarded. Discrepancy between values remains. | 8 | 7.44 | 0.56
X14 | 2010-04-27 13:50 | 362 | 357 | 1.3908 |  |  |  | 8 | 7.3 | 0.7
X14 | 2010-05-04 08:49 | 563 | 506 | 10.6642 |  |  |  | 8 | 7.53 | 0.47
X14 | 2010-05-11 08:20 | 501 | 471 | 6.1728 |  |  |  | 8 | 7.65 | 0.35
X14 | 2010-05-18 11:02 | 382 | 360 | 5.9299 |  |  |  | 7.8 | 8.02 | 0.22
X14 | 2010-05-25 08:40 | 229 | 243 | 5.9322 |  |  |  | 7.8 | 8.01 | 0.21
X14 | 2010-06-01 08:40 | 309 | 345 | 11.0092 |  |  |  | 7.3 | 8.14 | 0.84
X14 | 2010-06-08 08:36 | 376 | 380 | 1.0582 |  |  |  | 7.9 | 7.93 | 0.03
X14 | 2010-06-15 09:01 | 446 | 420 | 6.0046 |  |  |  | 8.1 | 7.87 | 0.23
X14 | 2010-06-22 09:56 | 454 | 438 | 3.5874 |  |  |  | 8.1 | 7.75 | 0.35
X14 | 2010-06-29 09:50 | 511 | 501 | 1.9763 |  |  |  | 7.95 | 7.35 | 0.6
X14 | 2010-07-06 09:56 | 383 | 395 | 3.0848 |  |  |  | 7.8 | 8.01 | 0.21
X14 | 2010-07-13 08:45 | 500 | 456 | 9.205 |  |  |  | 7.78 | 8.19 | 0.41
X14 | 2010-07-20 10:08 | 594 | 1693 | 96.1084 | Lab value correctly entered into emLine, but field value incorrectly entered. | Change Value | Field value changed; now correctly entered into emLine (see below). | 7.99 | 7.76 | 0.23
X14 | 2010-07-20 10:08 | 594 | 563 | 5.3587 | Lab value correctly entered into emLine, and incorrect field value amended (see above). | Let Value Stand | Discrepancy between lab and field values resolved. | 7.99 | 7.76 | 0.23
X14 | 2010-07-27 09:01 | 637 | 606 | 4.9879 |  |  |  | 7.99 | 8.11 | 0.12
X14 | 2010-08-03 13:14 | 311 | 338 | 8.3205 |  |  |  | 8.08 | 7.99 | 0.09
X14 | 2010-08-10 08:56 | 638 | 658 | 3.0864 |  |  |  | 7.78 | 7.92 | 0.14
X14 | 2010-08-17 08:50 | 660 | 614 | 7.2214 |  |  |  | 8.11 | 7.92 | 0.19
X14 | 2010-08-24 08:40 | 496 | 423 | 15.8868 |  |  |  | 8.09 | 8.67 | 0.58
X14 | 2010-08-31 14:10 | 610 | 650 | 6.3492 |  |  |  | 8.17 | 7.96 | 0.21
X14 | 2010-09-07 09:30 | 503 | 532 | 5.6039 |  |  |  | 7.93 | 8.27 | 0.34
X14 | 2010-10-21 11:30 | 471 | 426 | 10.0334 |  |  |  | 8.14 | 7.8 | 0.34
X14 | 2010-11-09 16:28 | 491 | 500 | 1.8163 |  |  |  | 8.19 | 7.89 | 0.3
X14 | 2010-11-18 10:30 | 1080 | 1010 | 6.6986 |  |  |  | 8.05 | 7.75 | 0.3
X14 | 2010-11-25 16:20 | 1020 | 1008 | 1.1834 |  |  |  | 7.73 | 8.21 | 0.48
X14 | 2010-12-02 08:42 | 984 | 1000 | 1.6129 |  |  |  | 7.95 | 7.7 | 0.25
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X2 | 2010-01-11 15:05 | 305 | 266 | 13.6602 |  |  |  | 7.5 | 6.94 | 0.56
X2 | 2010-02-22 12:01 | 322 | 288 | 11.1475 |  |  |  | 7.8 | 7.31 | 0.49
X2 | 2010-03-10 08:23 | 318 | 132 | 82.6667 | Field and lab values correctly entered into emLine. | Request Retest | Retest not possible because was requested too late. Remainder of sample had been discarded. | 7.8 | 7.27 | 0.53
X2 | 2010-04-14 15:07 | 352 | 290 | 19.3146 |  |  |  | 8.2 | 7.03 | 1.17 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X2 | 2010-05-03 16:13 | 113 | 188 | 49.8339 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.7 | 7.46 | 0.24
X2 | 2010-06-04 12:31 | 92 | 94 | 2.1505 |  |  |  | 7.6 | 7.81 | 0.21
X2 | 2010-07-08 15:10 | 139 | 131 | 5.9259 |  |  |  | 7.91 | 7.82 | 0.09
X2 | 2010-08-03 16:10 | 201 | 197 | 2.0101 |  |  |  | 8.16 | 7.68 | 0.48
X2 | 2010-09-02 09:43 | 208 | 189 | 9.5718 |  |  |  | 7.96 | 7.37 | 0.59
X2 | 2010-10-18 11:45 | 226 | 220 | 2.6906 |  |  |  | 7.96 | 7.6 | 0.36
X2 | 2010-11-09 14:04 | 268 | 258 | 3.8023 |  |  |  | 8.05 | 7.98 | 0.07
X2 | 2010-12-01 17:35 | 271 | 130 | 70.3242 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. Multiple flagged conductivity values on December 1, 2010 suggests calibration on field meter may have been off. | 7.92 | 7.2 | 0.72
X22b | 2010-01-11 11:05 | 1360 | 1146 | 17.079 |  |  |  | 7.5 | 6.4 | 1.1 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X22b | 2010-02-22 13:11 | 1290 | 1210 | 6.4 |  |  |  | 7.5 | 7.44 | 0.06
X22b | 2010-03-10 08:06 | 1270 | 1171 | 8.1114 |  |  |  | 7.6 | 7.04 | 0.56
X22b | 2010-04-13 13:36 | 1280 | 1230 | 3.9841 |  |  |  | 7.3 | 6.89 | 0.41
X22b | 2010-05-03 15:09 | 369 | 344 | 7.0126 |  |  |  | 7.4 | 7.23 | 0.17
X22b | 2010-06-04 17:45 | 1200 | 1178 | 1.8503 |  |  |  | 7.5 | 7.14 | 0.36
X22b | 2010-07-10 13:20 | 1230 | 1163 | 5.5997 |  |  |  | 7.43 | 7.33 | 0.1
X22b | 2010-08-03 10:30 | 1240 | 1215 | 2.0367 |  |  |  | 7.63 | 6.55 | 1.08 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X22b | 2010-09-01 11:41 | 1270 | 1222 | 3.8523 |  |  |  | 7.64 | 7.25 | 0.39
X22b | 2010-10-20 09:00 | 1310 | 1328 | 1.3647 |  |  |  | 7.67 | 6.25 | 1.42 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X22b | 2010-11-08 10:19 | 1340 | 1190 | 11.8577 |  |  |  | 7.69 | 7.46 | 0.23
X22b | 2010-12-01 11:58 | 1350 | 1340 | 0.74349 |  |  |  | 7.63 | 6.9 | 0.73
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X23 | 2010-01-12 12:00 | 8450 | 7990 | 5.5961 |  |  |  | 6.1 | 6.23 | 0.13
X23 | 2010-04-27 10:58 | 7030 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 2.9 | 2.96 | 0.06
X23 | 2010-05-06 08:05 | 7450 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 6.3 | 6.55 | 0.25
X23 | 2010-06-10 11:22 | 7920 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.9 | 6.99 | 1.09 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X23 | 2010-07-10 15:50 | 8090 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.9 | 6.16 | 0.26
X23 | 2010-08-05 10:57 | 7900 | 7450 | 5.8632 |  |  |  | 6.28 | 7.12 | 0.84
X23 | 2010-09-02 09:25 | 7900 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 6.32 | 6.18 | 0.14
X23 | 2010-10-07 13:17 | 7610 | 4300 | 55.5835 | Field and lab values correctly entered into emLine. | Let Value Stand | Retest request would be filed past hold time. Discrepancy between values remains. | 6.05 | 6.19 | 0.14
X23 | 2010-11-04 11:26 | 7740 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 6.81 | 6.61 | 0.2
X3 | 2010-01-11 14:55 | 303 | 266 | 13.0053 |  |  |  | 7.6 | 6.92 | 0.68
X3 | 2010-02-22 12:11 | 314 | 292 | 7.2607 |  |  |  | 7.9 | 7.35 | 0.55
X3 | 2010-03-09 12:37 | 316 | 260 | 19.4444 |  |  |  | 7.9 | 7.39 | 0.51
X3 | 2010-04-14 15:50 | 327 | 290 | 11.9935 |  |  |  | 8.2 | 7.02 | 1.18 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X3 | 2010-05-03 16:19 | 117 | 114 | 2.5974 |  |  |  | 7.4 | 7.74 | 0.34
X3 | 2010-06-04 11:50 | 84 | 86 | 2.3529 |  |  |  | 7.7 | 7.83 | 0.13
X3 | 2010-07-08 15:29 | 118 | 111 | 6.1135 |  |  |  | 7.85 | 7.91 | 0.06
X3 | 2010-08-03 15:41 | 189 | 183 | 3.2258 |  |  |  | 8.16 | 7.86 | 0.3
X3 | 2010-09-02 10:14 | 186 | 205 | 9.7187 |  |  |  | 7.99 | 7.34 | 0.65
X3 | 2010-10-21 13:40 | 234 | 225 | 3.9216 |  |  |  | 8.05 | 7.57 | 0.48
X3 | 2010-11-09 14:42 | 258 | 256 | 0.77821 |  |  |  | 8.1 | 7.92 | 0.18
X3 | 2010-12-01 17:08 | 259 | 230 | 11.8609 |  |  |  | 7.95 | 7.62 | 0.33
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X3A | 2010-09-02 11:00 | 206 | 192 | 7.0352 |  |  |  | 7.99 | 7.5 | 0.49
X3A | 2010-10-21 09:35 | 244 | 268 | 9.375 |  |  |  | 8.11 | 7.84 | 0.27
X3A | 2010-11-09 15:08 | 264 | 266 | 0.75472 |  |  |  | 8.14 | 7.97 | 0.17
X3A | 2010-12-01 17:28 | 263 | 240 | 9.1451 |  |  |  | 7.99 | 7.75 | 0.24
X4 | 2010-01-11 11:49 | 2210 | 1791 | 20.9448 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 3.4 | 3.02 | 0.38
X4 | 2010-02-22 14:06 | 1970 | 1820 | 7.9156 |  |  |  | 4 | 4.82 | 0.82
X4 | 2010-03-09 11:52 | 1390 | 1350 | 2.9197 |  |  |  | 6.5 | 6.35 | 0.15
X4 | 2010-04-13 10:56 | 1340 | 1350 | 0.74349 |  |  |  | 6.1 | 6.36 | 0.26
X4 | 2010-05-03 15:39 | 1010 | 877 | 14.0964 |  |  |  | 4.7 | 4.51 | 0.19
X4 | 2010-06-04 08:58 | 1940 | 1868 | 3.7815 |  |  |  | 3.2 | 3.12 | 0.08
X4 | 2010-07-10 15:35 | 1930 | 1871 | 3.1044 |  |  |  | 3.32 | 3.12 | 0.2
X4 | 2010-08-03 14:00 | 1950 | 2020 | 3.5264 |  |  |  | 3.39 | 460 | 456.61 | Lab value correctly entered into emLine, but field value incorrectly entered. | Change Value | Field value changed; now correctly entered into emLine (see below).
X4 | 2010-08-03 14:00 | 1950 | 2020 | 3.5264 |  |  |  | 3.39 | 4.6 | 1.21 | Lab value correctly entered into emLine, and incorrect field value amended (see above). | Let Value Stand | Large discrepancy between lab and field values resolved, but small discrepancy between values still remains.
X4 | 2010-09-02 12:27 | 1880 | 1752 | 7.0485 |  |  |  | 3.48 | 4.69 | 1.21 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X4 | 2010-10-20 11:20 | 1590 | 1511 | 5.0951 |  |  |  | 5.61 | 6.65 | 1.04 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X4 | 2010-11-08 12:05 | 1860 | 1632 | 13.0584 |  |  |  | 3.73 | 4.49 | 0.76
X4 | 2010-12-01 08:24 | 1310 | 1429 | 8.6893 |  |  |  | 6.32 | 6.36 | 0.04
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X5 | 2010-03-22 17:51 | 1810 | 1580 | 13.5693 |  |  |  | 8.1 | 7.11 | 0.99
X5 | 2010-03-30 10:41 | 1800 | 1700 | 5.7143 |  |  |  | 7.9 | 7.02 | 0.88
X5 | 2010-04-06 08:36 | 1880 | 1720 | 8.8889 |  |  |  | 7.8 | 7 | 0.8
X5 | 2010-04-13 09:10 | 1840 | 1561 | 16.4069 |  |  |  | 7.6 | 7.41 | 0.19
X5 | 2010-04-20 14:36 | 1820 | 1947 | 6.7428 |  |  |  | 7.7 | 6.85 | 0.85
X5 | 2010-04-27 13:56 | 1450 | 1399 | 3.5802 |  |  |  | 7.9 | 7.19 | 0.71
X5 | 2010-05-04 08:37 | 1620 | 1433 | 12.2502 |  |  |  | 7.9 | 7.06 | 0.84
X5 | 2010-05-11 08:05 | 1490 | 1390 | 6.9444 |  |  |  | 7.9 | 7.45 | 0.45
X5 | 2010-05-18 10:45 | 1350 | 1310 | 3.0075 |  |  |  | 8 | 7.37 | 0.63
X5 | 2010-05-25 08:30 | 1430 | 1441 | 0.76628 |  |  |  | 7.9 | 7.51 | 0.39
X5 | 2010-06-01 08:15 | 1670 | 1740 | 4.1056 |  |  |  | 8.1 | 7.51 | 0.59
X5 | 2010-06-08 08:26 | 1740 | 1664 | 4.4653 |  |  |  | 7.9 | 7.49 | 0.41
X5 | 2010-06-15 08:28 | 1760 | 1670 | 5.2478 |  |  |  | 8.2 | 7.48 | 0.72
X5 | 2010-06-22 09:43 | 1710 | 1605 | 6.3348 |  |  |  | 8.2 | 7.63 | 0.57
X5 | 2010-06-29 09:42 | 1770 | 1725 | 2.5751 |  |  |  | 7.82 | 6.8 | 1.02 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains.
X5 | 2010-07-06 09:43 | 2280 | 1705 | 28.8582 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.89 | 7.57 | 0.32
X5 | 2010-07-13 08:32 | 1810 | 1567 | 14.3915 |  |  |  | 7.97 | 7.85 | 0.12
X5 | 2010-07-20 09:51 | 1830 | 1674 | 8.9041 |  |  |  | 7.85 | 7.69 | 0.16
X5 | 2010-07-27 08:35 | 1780 | 1589 | 11.3387 |  |  |  | 7.92 | 7.83 | 0.09
X5 | 2010-08-10 08:32 | 1780 | 1806 | 1.4501 |  |  |  | 7.76 | 8.02 | 0.26
X5 | 2010-08-17 08:40 | 1750 | 1594 | 9.3301 |  |  |  | 7.99 | 7.74 | 0.25
X5 | 2010-08-24 09:00 | 1740 | 1266 | 31.5369 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 8.01 | 7.98 | 0.03
X5 | 2010-08-31 14:45 | 1770 | 1860 | 4.9587 |  |  |  | 7.84 | 7.64 | 0.2
X5 | 2010-09-07 09:45 | 1840 | 1708 | 7.4408 |  |  |  | 7.84 | 7.65 | 0.19
X5 | 2010-11-18 10:40 | 2360 | 2219 | 6.1585 |  |  |  | 7.94 | 7.3 | 0.64
X5 | 2010-11-25 16:37 | 2500 | 2414 | 3.5002 |  |  |  | 7.43 | 7.57 | 0.14
X5 | 2010-12-02 09:09 | 2470 | 2590 | 4.7431 |  |  |  | 7.73 | 6.88 | 0.85
X5P | 2010-01-11 12:12 | 1860 | 1559 | 17.6075 |  |  |  | 7.6 | 6.63 | 0.97
X5P | 2010-02-22 08:50 | 1740 | 1569 | 10.3354 |  |  |  | 7.7 | 7.21 | 0.49
X5P | 2010-03-02 11:30 | 1750 | 1420 | 20.8202 | Field and lab values correctly entered into emLine. | Let Value Stand | Discrepancy between lab and field values remains. | 7.7 | 7.03 | 0.67
X5P | 2010-04-13 11:00 | 1780 | 1488 | 17.8703 |  |  |  | 7.5 | 7.84 | 0.34
X5P | 2010-05-03 15:44 | 877 | 744 | 16.4096 |  |  |  | 7.8 | 7.75 | 0.05
X5P | 2010-06-04 09:24 | 1740 | 1682 | 3.3898 |  |  |  | 7.7 | 7.13 | 0.57
X5P | 2010-07-10 15:12 | 1780 | 1680 | 5.7803 |  |  |  | 7.79 | 7.64 | 0.15
X5P | 2010-08-03 13:47 | 1800 | 1665 | 7.7922 |  |  |  | 8 | 7.17 | 0.83
X5P | 2010-09-02 12:12 | 1760 | 1687 | 4.2356 |  |  |  | 7.75 | 7.53 | 0.22
X5P | 2010-10-20 10:55 | 2070 | 1955 | 5.7143 |  |  |  | 8.06 | 7.38 | 0.68
X5P | 2010-11-08 11:55 | 2150 | 1918 | 11.4061 |  |  |  | 7.92 | 7.84 | 0.08
X5P | 2010-12-01 08:50 | 2130 | 2204 | 3.4149 |  |  |  | 7.95 | 7.45 | 0.5
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit
X7 | 2010-06-01 16:20 | 8550 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.4 | 5.59 | 0.19
X7 | 2010-10-06 13:42 | 8490 | >3999 | #VALUE! | Field and lab values not comparable because conductivity greater than maximum measurable level for field meter used. |  |  | 5.18 | 5.42 | 0.24
RPD > 50% or pH difference > 1.5 pH unit
RPD > 20% or pH difference > 1 pH unit

## Refs
Change Value
Let Value Stand
Remove Value
Request Retest